## Sheet1
Alarm Condition | Condition Detected | Response
PARAM NOT CONFIRMED | Control parameter modified, but not confirmed | Previous parameters maintained, "Parameters Not Confirmed" message displayed, cleared by confirming or dismissing parameter changes
LOW PRESSURE | Pressure lower than expected (possible airway disconnect) | Maintain operation, sound auditory alarm and flash screen at full brightness, cleared with onscreen confirmation
HIGH PRESSURE | Pressure above tolerable threshold (when Ambu-bag pressure release valve opens) | Maintain operation, sound auditory alarm and flash screen at full brightness, cleared with onscreen confirmation
POSITION ERROR | Stepper driver error or unexpected position/velocity value reported (possible sensor fault or mechanical interferenct) | Trigger Rehoming procedure on next Expiratory Pause, if error persists, sound auditory alarm and flash screen at full brightness, cleared with onscreen confirmation and successful Rehoming
SYSTEM FAULT | Fault/error/erroneous behavior detected in controller, stepper driver or sensors (possible electrical fault, overheat, etc.) | "System Fault" message displayed with specific fault information, sound auditory alarm and flash screen at full brightness, cleared with onscreen confirmation and correction of fault
HOMING FAULT | Homing or Rehoming procedure fails (i.e. unexpected limit switch or position values) | Re-attempt Homing procedure, "Homing Fault" message displayed upon subsequent failure, cleared with onscreen confirmation and successful Rehoming